## Sheet1
Parameter | Default | Unit | Description
[PartView]
AngleRadiusPct | 50 | Percent | Radius of the arc showing an angle between lines as a percent of the distance from start to end of shorter parent line
DisplayDiagnostics | 0 | 0=False/1=True | Display onscreen part view diagnostics
ShowPatternFeature | 0 | 0=False/1=True | Show/Hide the Pattern Feature graphic in the Part View
DisplayDXFOverlays | 0 | 0=False/1=True | Show/Hide the DXF Overlay in the Part View
UseDifferentRunColors | 0 | 0=False/1=True | Allow disabling of the color display of features in the partview, when they are displayed as multiple part playback results. When enabled all features will be black for Passed or Red for Failed.
NumCircleSegments | 64 | # of segments | Ability to specify the minimum number of segments used for drawing circles and arc features in the partview.
FOVRadius | 0 | *
MinDistForPan | 15 | *
MinDistForZoom | 15 | *
MinDistForGoodSelectMarquee | 10 | *
[Beta]
ScrimDialogs | 1 | 0=False/1=True | If 1 then displays the backgroung scrim when a dialog is displayed
DelayForUserToRead | 0 | * | If 1 then the various "Please wait..." prompts will not be displayed during program execution (would improve throughput)
ShortCuts | 1 | 0=False/1=True | Enable/Disable keyboard shortcuts
[Sounds]
TouchProbe | 1 | 0=Off/1=On | Substitute the touchprobe point entry sound. Add parameter, and add a new wave file called touchpoint.wav to the M3 running directory to enable the new touch probe sound.
[TouchProbe]
TeachArtifact | Sphere | Sphere=Cal Sphere / RingGage=Ring Gage | Sets the Touch Probe calibration artifact type (depracated)
TouchProbeVelocityCompMicroSeconds | 200 | Ms | Used to compensate for propagation delay of the touch probe circuit
[ProbePath]
ParkEnabled | 0 | 0=False/1=True | Set the park position machine coordinates in the Joystick settings screen, and set the enable flag to activate.
[Print]
PrintEmpty | 0 | 0=False/1=True | If 1 then will print even if it would otherwise be empty (e.g. a part view print with no features)
FontSize | 18 | * | The font size of non-annotation printed text
AnnotationFontSize | 16 | * | The font size of annotation printed text
LeftMargin | 40 | * | Sets printout left margin size
RightMargin | 20 | * | Sets printout right margin size
TopMargin | 40 | * | Sets printout top margin size
BottomMargin | 40 | * | Sets printout bottom margin size
MinAnnotationScaling | 100 | % | Sets the minimum scaling percentage that can be used for annotaion printing
[Options]
GarantOption | 0 | 0=False/1=True | Enables the Garant OEM skin
CAA3 | 0 | 0=False/1=True | Enables Schneider corrections
ShapeCal | 0 | 0=False/1=True | Enables the Shape Calibration function/Displays the shape calibration setup screen
AutoLaunchLog | 0 | 0=False/1=True | Causes the mlxlog.exe program to launch each time the Mx is launched
NelaOption | 0 | 0=False/1=True | Enables Nela logo in to toolbar
DisplayB | 0 | 0=False/1=True | Changes the Q axis DRO label to B
FeedbackButton | 0 | 0=False/1=True | Enables Feedback Application to run when long press on Mx button is executed
MotionReturnToVideo | 1 | 0=False/1=True | Enables video motion return to Live Video function(when FOVPlayback=1)
DemoStage | 0 | 0=False/1=True | Enables FOV playback mode
FOVPlayback | 0 | 0=False/1=True | Enables FOV playback mode
DemoStageTp | 0 | 0=False/1=True | Shows demo tp hit button in top toolbar
Demo | 0 | 0=Red, 1= Grey | Convert all Red color Error dialog to a grey color.
PurePartFiles | 0 | 0=False/1=True | Enables .mxlpure that remain unconverted after playback
RunPartFromBarcode | 0 | 0=False/1=True | Barcode string entered that matches an existing mlxpart2 file in the parts directory, will open the matching program and run it
PatternDatumTopLevel | 1 | 0=False/1=True | Displays the "Measure Registration Pattern" button on the top level of the main measure toolbar.
LargeInsulationLabels | 0 | 0=False/1=True | Enables alternate coefficient labels for Insulation results. When enabled will display Min/Max/Avg/Con
EncryptNLEC | 1 | 0=False/1=True | Set to 0 to export un-encrypted
DualBlobInsulation | 0 | 0=False/1=True | Allow the collection of Insulation Data one layer at a time(Insulation Module Required). Enable the parameter, then use the keyboard "Ctrl" key to collect one insuilation layer per mouse click.
StaticPixel | 0 | 0=False/1=True | Ball End Mill inspection functionality to capture pixels statically within the FOV. Can be configured to capture either the "light" or "dark" pixels of the current FOV. Supporting the inspection of Ball End Mill wear.
ProbeRackWarnWindowDisable | 0 | 0=False/1=True | Used to disable the probe change dialig box when using a probe rack.
[Startup]
Probe | sense | Sense / Xhair / Active | Sets the startup probe to MeasureLogic/Xhair/Active
ZoomAll | 0 | 0=False/1=True | Sets the starup video zoom type to Zoom All(Fill)
Zoom1to1 | 0 | 0=False/1=True | Sets the starup video zoom type to Zoom 1:1
ZoomFill | 1 | 0=False/1=True | Sets the starup video zoom type to Zoom Fill
[ShapeCal]
HideButton | 0 | 0=False/1=True | When Shape Cal is enabled, allows the shapecal button to be displayed or hidden
StartupMode | 0 | MagsShapeCalNone-coal | Sets startup shapecal button to None
" |  | MagsShapeCalRound-coal | Sets startup shapecal button to Round
" |  | MagsShapeCalSquare-coal | Sets startup shapecal button to Square
OverrideMode |  | MagsShapeCalNone-coal | Forces the shape cal to stay at the None mode
" |  | MagsShapeCalRound-coal | Forces the shape cal to stay at the Round mode
" |  | MagsShapeCalSquare-coal | Forces the shape cal to stay at the Square mode
" |  | MagsShapeCalFlat-coal | Forces the shape cal to stay at the Flat mode
EdgeOff1 | 0 | MM | The X and Y offset when in ShapeCalNone mode
[Hardware]
ReportScaleErrors | 1 | 0=False/1=True | Enables/Disables the reporting of scale errors
ZeroOnInputs | 0 | * | If 1, causes input lines 1-3 to act like a Dro Zero for that axis
SchneiderTrigger | 1 | * | If 1, causes input lines 2 to start a program run
GoAwayAfterSecs | 0 | Seconds | Amount of time the "connection" messages are displayed (0 is wait-for-click, -1 shuts them off)
XAxis | 0 | Hardware Axis Channel | Allows remapping of the HW axis Channel to the top level DRO display(0=X, 1=Y, 2=Z, 3=Q)
YAxis | 1
ZAxis | 2
QAxis | 3
FreezeImageOnZoom | 0 | 0=False/1=True | Freezes live video image during motorized zoom indexing(Opto Turret)
AutoClearHomeCompleteMessage | 0 | Seconds | Seconds to display message. 0=Require click to clear.
[Laser]
UseDedicatedLine | 0 | 0=False/1=True | If 1, laser output is via a dedicated laser output, otherwise it uses a normal output line
BackwardCompatible | -1 | * | If a laser output uses a normal output line, a -1 uses DtoA 6, else the output line specified
For MicroEpsilon Laser
QtyPointsToAvg | 300 |  | Number of raw laser points to average when creating a point.
MeasuringRate | 5000 |  | Laser measuring rate (in Hz). If the laser is incapable of the requested rate, the next closest rate (rounding up) is selected. If the requested rate is above the max rate for the laser, the max rate is selected.
ProbingSpeed | 100 |  | Percentage of max velocity at which the CNC moves when the laser is taking a measurement (when part following is not enabled).
NonProbingSpeed | 100 |  | Percentage of max velocity at which the CNC moves during laser operation when not acquiring data (i.e. approach speed when traveling from VED to laser position, gotos between scan lines in a patch probe).
PulseDistance | 10 |  | Raw laser point spacing (in microns)
RangePercent | 90 |  | Percent of the laser range (centered around the mid) that defines being "in focus".
FocusSeekUp | 10 |  | Max Z distance (in mm) that the laser will travel up from the current position to find range when seeking focus.
FocusSeekDown | 5 |  | Max Z distance (in mm) that the laser will travel down from the current position to find range when seeking focus.
AllowTransition | 0 |  | Enables laser transition probe (for use in cross calibration).
AllowRepeat | 0 |  | Enabled the use of Point-->MultiPoint measure mode to return source laser scan points directly to the feature list.
PlaneClipAngle | 15 |  | Angle snap (in degrees) for laser patch probe layout.
AllowPoint | 1 |  | Enable/disable laser point probe.
AllowLine | 1 |  | Enable/disable laser line probe.
AllowRect | 1 |  | Enable/disable laser patch probe.
AutoZVelPer | 20 |  | Percentage of max velocity at which the CNC moves when part following (Auto Z Pos) is enabled.
ReposCycleTime | 10 |  | Time interval (in ms) at which to update the Z position of the laser for optimal measuring distance during part following.
[Features]
MaxNameLength | 10 | 1-255 | Sets number of characters in a feature name.
[Translation]
AllowChangeEnglish | 0 | 0=False/1=True | If 1, allows the user to translate the english text file
[System]
SeriousErrorsOnStartup | 0 | * | Fov file and nlec file warning at startup if set to 1
UpdateEverySecs | 0.03 | Seconds | Main loop time (desired - actual may vary)
BackupEveryDays | 7 | Days | Sets the frequency of the Mx Auto Backup function
DisplayMainLogError | 1 | * | Determines if the top toolbar is displayed in red when an exception occurs
[NLEC]
ReportFile | 1 | 0=False/1=True | Enables output of the NLEC calibration results text file
ShowTeach | 1 | 0=False/1=True | Set to "1" to display button
DualGrid | 0 | 0=False/1=True | For dual level NLEC
[Desktop]
ShowStitchButton | 1 | 0=False/1=True | Show/Hide the stitch button in the top toolbar
LongPressFeedbackWaitSecs | 1 | Seconds | Amount of time to wait (after pressing) before the Feedback button is triggered
LongPressWaitSecs | 1 | Seconds | Amount of time to wait (after pressing) before most long-button presses are activated
ShowAll3dButtons | 0 | 0=False/1=True | Displays 3D feature measurement buttons in the top level of the measure toolbar
OskPopUp | 1 | 1=Display Keyboard, 0=Do not display keyboard | Supress automatic display of the Windows On-Screen Keyboard, when text entry fields are active.
[Measure]
Allow3dCircle | 0 | 0=False/1=True | Circle construction from a cone and a plane.
SymmetryLineRule | 1 | 0=False/1=True | Provides backward compatibility for V1 symmetry line handling
UseConeLengthTest | 0 | 0=False/1=True | Set to "0" to bypass the test
ProbeSmallAngle | 0 | 0=Off/1=On | Enable the direct probing of angle features near 180 degrees. Enable the following parameter to lift the near 180 restriciton on directly probed angle features.
MaxParents | 250 | Parents | Constructions from parent populations of more than 250 features
GearFilter | 1 | 1=Default | Added a new Gear Result Filtration Parameter to the metlogix.ini file. Setting the following filter applies a results filter to the Gear Pin result based on a standard deviation from the best fit master circle result.
GearFilterParams | 2.5 | Unit is STD DEV
[Programming]
UseLastUserMessageResponse | 1 | * | 1=display last entered user text, 0=display each user message prompt step as blank in program playback.
RequireUserMessageResponse | 0 | 0=Disabled 1=Enabled | Require a User Response in User Message Ask steps in program playback. When enabled, a program cannot advance past a User Message Ask step unless some string has been entered for it.
ClearStitchOverride | 0 | 0=Disabled 1=Enabled | Lock the last stitched image as displayed in live Video, after program playback.
NewPartBeforeOpen | 0 | 0=Disabled 1=Enabled | Forces a New Part(clear), prior to loading a part program. This mechanism can improve loading performance in very large part programs(15K+ features).
ClearTeachOnNewPart | 1 | 0=Off/1=On | Clear existing Advanced Edge Teach values on execution of a new part. Added metlogix.ini parameter to disable this behavior.
FocusLockZone | 0.05 | MM | The zone in which a focus lock stays active
TargetArrowLenInPixels | 100 | Pixels | The length of the program playback guidance arrow
ShowAnimatedDisplay | 0 | 0=False/1=True | If 1, causes the auto program part view display to be animated during operation
UseZPlaybackNavigation | 1 | 0=False/1=True | Includes/excludes Z position in calculation of playback position
MarkupWait | 1 | Seconds | The time that the system waits for settling, during programmatic markup prints / archives
AllowQuickSave | 0 | 0=False/1=True | Enables quick save function(long press on program file name at feature list header)
PayAttentionToStitchAtRun | 1 | * | If 1, allows a stitch to stay active when a program is run
RunManyOfTheSame | 0 | 0=False/1=True | Sets "Run Many of the Same" function as the Program Default
ForceCartMode | 1 | 0=False/1=True | Enables data output in the currently selected coordinate system mode(cartesian or polar)
ShowLoop | 0 | 0=False/1=True | Enables Loop Part Program Exhibition Feature(via propgram properties item)
MaxPalletSize | 100 | # of pallet pockets | Sets the maximum number of pallet pockets that can be used
PowerAssist | 1 | 0=False/1=True | Enables/Disables Xhair power assist mode for CNC systems
MaxFormBeforeRecovery | 1 | * | maximum form for a successful probed feature during program playback
RemeasureFromExit | 0 | 0=False/1=True | Enables/Disables capability to update measurement results from Exit mode via remeasure edit function
RecordFocusStep | 1 | * | If 0, disables recording of a focus step
UseFixtureNames | 0 | 0=False/1=True | Enables programs to share a common permanent fixture reference. Add fixture name in Program Properties in edit mode when enabled
PermanentFixture | 0 | 0=Normal/1=Default to Permanent | Controls default state of Permanent Fixture mode
SamePositionPrompt | 1 | 1=Normally enabled prompt 0=Disables prompt and assumes "Yes" (part in same position) 2=Disables prompt and assumes "No"(part has changed positions) | Allow the automatic selection of a response to the temporary fixture dialog before running a program
DisplayPrint | 0 | 0=Off/1=On | Parameter that enables the ability to generate a printout of a part programs steps. Enable the ini parameter below to display the new print button in the bottom toolbar in edit mode.
FreezeDuringMove | 0 | 0=False/1=True | Set to 1 to enable. 0, or off by default
FailedPalletZClearance | 9999 | MM | Sets the distance(in the Z axis positive direction) that the system will clear to when advancing a pocket via the "Continue pallet on failure" option
WriteBuffer | 20 | Megabytes | Manual configuration of file write buffer size. When extermely large patterns are being used (on stitched images) it may be necessary to increase size of buffer to support successful program conversion.
ClearTeachOnNewPart | 1 | 0=False/1=True | Added new functionality to clear existing Advanced Edge Teach values on execution of a new part. Added a metlogix.ini parameter to disable this behavior.
IgnorePocketPrompt | 0 | 0=False/1=True | Set to "1" to enable pocket selection dialog override
ShowCurRunStep | 0 | 0=False/1=True | Enables program playback timer, and program feature counter to playback, displayed in the program status toolbar.
FreezeOnFovPlayback | 0 | 0=False/1=True | Changed the default behavior for programs set to "FOV Playback Mode". The image is no longer frozen at the start of program by default. Parameter can be used to restore the old behavior
ShowCurRunStep | 1 | 0=False/1=True | Added new program playback timer, and program feature counter to playback, displayed in the program status toolbar
[Display]
RunNextToEnter | 0 | 0=False/1=True | Displays a program run button in the bottom right toolbar
ClipQto360 | 0 | 0=-180 to 180deg/1=0-360deg | Sets the Q axis encoder DRO format to either -180 to 180, or to 0 to 360
ComboBoxPopups | 1 | 0=False/1=True | If 0, disables popups in settings and forces "cycling" behavior
ToolTips | 0 | 0=False/1=True | Enable/Disable "hover" tooltip display
UndoNextToEnter | 0 | 0=False/1=True | Displays a undo button in the bottom right toolbar
[Stitch]
MaxStitches | 12 | Frames | Sets the number of frames(stitches) that can be collected for a superimage
PresetColor0 | 0, 255, 0 | RGB | Sets the color of the stitch preset button 1
PresetColor1 | 0, 0, 255 | RGB | Sets the color of the stitch preset button 2
PresetColor2 | 255, 140, 0 | RGB | Sets the color of the stitch preset button 3
[Overlay]
Thickness | 1 | Pixels | Thickness of markup leader lines
X | 0 | Pixels | Small adjustment to better align live video overlays with the actual image
Y | 0 | Pixels | Small adjustment to better align live video overlays with the actual image
Width | 0 | Pixels | Small adjustment to better align live video overlays with the actual image
Height | 0 | Pixels | Small adjustment to better align live video overlays with the actual image
[RS232]
IncludeLabels | 0 | * | *
EndChars1 | 13 | * | Post form chars
EndChars2 | 10 | * | Post form chars
EndChars3 | -1 | * | Post form chars
EndChars4 | -1 | * | Post form chars
EndChars5 | -1 | * | Post form chars
StartChars1 | -1 | * | Pre form chars
StartChars2 | -1 | * | Pre form chars
StartChars3 | -1 | * | Pre form chars
StartChars4 | -1 | * | Pre form chars
StartChars5 | -1 | * | Pre form chars
SeparatorChars1 | 32 | * | Separator char added between multiple fields in a single line
SeparatorChars2 | -1 | * | Separator char added between multiple fields in a single line
SeparatorChars3 | -1 | * | Separator char added between multiple fields in a single line
SeparatorChars4 | -1 | * | Separator char added between multiple fields in a single line
SeparatorChars5 | -1 | * | Separator char added between multiple fields in a single line
LabelChars1 | 32 | * | Separator char added between a field label and value
LabelChars2 | -1 | * | Separator char added between a field label and value
LabelChars3 | -1 | * | Separator char added between a field label and value
LabelChars4. | -1 | * | Separator char added between a field label and value
LabelChars5 | -1 | * | Separator char added between a field label and value
PortId | 1 | Com Port | Designate Serial Port ID number for RS232 Output
Baud |  | 9600 | Designate Baud Rate for RS232 Output
Parity |  | 0 | Designate Parity for RS232 Output
ByteSize |  | 8 | Designate Byte Size for RS232 Output
StopBits |  | 0 | Designate number of stop bits for RS232 Output
CtsFlow |  | 0 | Designate Cts Flow for RS232 Output
DsrFlow |  | 0 | Designate Dsr flow for RS232 Output
ToClipboard | 0 | 0=False/1=True | Support RS232 export to clipboard functionality. When enabled all RS232 Send strings will go to the clipboard.
LastSelectionSet | 0 | 0 = Always display dialog with nothing selected 1 = Remember the last selection set (Default) 2 = Always display dialog with everything selected
[Cnc]
AddLightStepForEveryGoto | 1 | * | Enable/Disable addition of a light control step for each goto recorded in a program
WarnOnMove | 1 | 0=False/1=True | When Disabled the CNC Warning Dialog for playback is not displayed
ActiveJoystickButtons | 0 | 0=False/1=True | When enabled the onscreen joystick status buttons will function as command output buttons, for setting Axis Lock or Speed Modes.(Only Supported in 3 Button Joystick Configurations)
[OEX]
StartupMode | OEXRound/OEXSquare | * | OEXRound=Enables round schneider oex correction at startup / OEXSquare=Enables square oex schneider correction at startup
OutputLine | 3 | * | Which output line to use to set schneider oex type
OutputHigh | 1 | * | The level of that output line
[DxfImport]
ProximityThreshold | 0.25 | MM | Specifies the matching criteria for the Import Nominals from DXF function (How close the dxf nominals are to the measured actuasl to be considered a match)
UseSystemUnits | 0 | 0=False/1=True | Sets whether dxf import uses the system units or the dxf file units
ContiguityThreshold | 0.0001 | MM | Sets how close converging dxf entities need to be to be considered a "closed loop" in the Overlay (profile) function.
[Security]
NumUsers | 5 | User Login Accounts | Sets the number of User type accounts available in Security Setup
[Lights]
Comm Port | 2 | Com Port | Sets light control Com Port for RS232 light control
RingMove | 0 | 0=False/1=True | Enables the Ztrack Ring Light Handle
RingUIRange | 250 | * | Sets the range for the Ztrack Ring Light Slider in the User Interface
ZTracChannel | -1 | Light Control Channel Number | Sets the channel to be used for Z-track
ZTracQMaxDistance | 0 | * | Sets the range for the Ztrack Ring Light handle in the User Interface
ZTracMax | 1000 | * | *
ShutterSpeed | 0 | 0=False/1=True | Enables the Exposure Control slider in the light control interface for supported cameras
DisplayLightCal | 0 | 0=False/1=True | Enables the restore calibrated light levels button in the light control interface
Output0Or1 | 0 | * | If 1, causes light outputs to go from low to high, instead of a range using the DtoA
DualLightControl | 0 | 0=Single Map 1=Dual Map | Support for (2) sets of light control mappings, based on the selected mag. This allows support for two sets of illuminators, that will share the same light channel number assignments.
InvertAllScales | 0 | 0 = 0V --> 5V Ramp / 1 = 5V --> 0V Ramp | Inverts the voltage output of the Light Control signals. When set to "1", the output becomes 5V --> 0V, when ramping the sliders from 0 to 100.
IndividualSliders | 0 | 0=False/1=True | Added new light control slider User Interface, utilizing 8 discrete slider bars. This option is enabled via the following metlogix.ini parameter.
[Flick]
Time1 | 1000 | Milliseconds | Amount of time to allow stage to travel before decelerating
Time2 | 1000 | Milliseconds | Amount of time to take to decelerate
Vmin | 2 | MM per second | Minimum travel velocity
Vmax | 20 | MM per second | Maximum travel velocity
Vfactor | 5 | Scalar | Amplifies the calculated flick velocity
[FocusCnc]
MaxRetract | 3 | MM | Sets the Focus Retract Range Maximum(Not Used*)
MinRetract | 0.5 | 33 | Sets the Focus Retract Range Minimum(Not Used*)
MaxVelocity | 1.5 | MM/Sec | Sets the Focus Maximum Velocity(Not Used*)
MinVelocity | 0.1 | MM/Sec | Sets the Focus Minimum Velocity(Not Used*)
maxScans | 8 | AF pass(both directions) | Sets the maximum number of scans before timeout(error displayed)
PosFactor | 0 | 0.0 - 1.0 scalar | In free run focus, position may need to be be shifted by direction to get contrast curves to fully overlap (expert diagnosis needed)
[Measuring]
MaxPoints | 1000 | Measurement points (0=off) | Sets the maximum number of points that can be collected for a measurement
MaxPointsBlob | 0 | Measurement points (0=off) | Sets the maximum number of points that can be collected for a blob measurement
MinimumPointsForFilter | 0 | Number of points | Minimum pts before a point filtration will occur
[Tolerance]
UseTolWarnings | 0 | 0=False/1=True | Enables Warning Display Color(Orange) as new deviation category.
TolWarningThreshold | 0.66 | Decimal percentage (of deviation) | Sets the threshold for displaying warning tolerance category for deviation from nominal
AllowNominalDontCare=1 | 0 | 0=False/1=True | When utilized, the nominal entry field for either X or Y can be left blank, whereby the tolerance will only control for position in 1 axis.
[AutoProgram]
MinNoPts | 8 | Number of points | The minimum number of points that will be created during an auto program
[CustomColor]
Enabled | 0 | 0=False/1=True | Allows changing the color of the top and bottom toolbars in the main M3 User Interface
TopR | 0 | 0 to 255 | Sets the Red value
TopG | 161 | 0 to 255 | Sets the Green value
TopB | 225 | 0 to 255 | Sets the Blue value
[Export]
IncludeQuotes | 1 | 0=False/1=True | Enable or Disable Quote Character separator in CSV output file
IncludeIJK | 1 | * | Enable or disable output of IJK info with export
TabDelimited | 0 | 0=Comma 1=Tab 2=Semicolon | Enable Tab Delimited separator format
UnicodeHeader | 0 | * | If 1, causes unicode files to be created for exports
IncludeDateTimeRow | 0 | 0=False/1=True | Include Date/Time row in export file
ShowIgs | 1 | 0=False/1=True | Enables support for .IGES format export of point data. When enabled the new export format will be available in the export menu.
IncludeDashes | 1 | 0=False/1=True | Filter the "-" and "+" characters out of the .txt output file
[Focus]
MinPoints | 24 | Scalar | Number of points needed for the focus parabolic fit routine
PercentThreshold | 65 | % | In focusing, data below 65% of max contrast is discarded
NoiseFloor | 100000 | Video code contrast units | This contrast (calculated by video code) is the minimum required
MoveThreshold | 20 | Encoder counts | In manual focus, the user must move Z this many counts to acquire a new data point
MaxPoints | 200 | Scalar | 200 valid points probably means data is mostly on one side of max contrast (in Z position)
[Video]
DxDeviceId | 1 | Direct Show Device Index Number | Specifies the Index number of the Direct Show Device to be laoded when the Direct Show Capture Source is used
RightMouseRotate | Motion | 0=False/1=True | Enables image rotation with right mouse longpress in Measure from Bitmap mode
RightMouseXlate | 0 | 0=False/1=True | Enables Image translation with right mouse longpress in Measure from Bitmap mode
RightMouseFinish | 0 | 0=False/1=True | When 1, suppresses motion of demo stage using right mouse
gradientError | 0.1 | Percentage of error | When playing back ved edges, amount of allowable error on matching edges before refing occurs
DustFilterAngle | 0 | Degrees | If greater than zero on circle and line firing, will throw out edges on dust that exceed the degrees specified
MaxExposure | 0.017 | Exposure Time in Ms | Sets the Max Exposure Value for Imaging Source cameras
MaxExposureMs | 1000 | Exposure Time in Ms | Sets the Max Exposure Value for uEye and Sentech GigE cameras
FOVErrorScale | 25 | Scalar | Grid cal error display exageration factor
PointCloudSizeMax | 10000 | Count | Max number of blob points to post to a measurement
MotionMinSum | 100 | * | How large the value changes must be for the system to consider motion to have occured
MotionThreshold | 25 | Count | Minimum difference between histograms to be in motion
MotionPad | 0 | Count | Width of a histogram peak
MotionFloor | 0.25 | % | Defines the floor of the bell curve of the largest gray pixel region
MotionTimeout | 2000 | Ms | Time to wait after motion has occurred to execute pattern rec in Auto Run mode
MotionLights | 0 | 0=False/1=True | Enable Disable onscreen diagnostic lights for state machine within Auto Play function
ImageDir | Default Metlogix/Images | System Path | Sets the source directory for program finder pattern images
PartDir | Default Metlogix/Parts | System Path | Sets the source directory for program finder part files
UseSquarePixelsToDraw | 0 | 0=False/1=True | Enable to force DXF Overlay draw to assume square pixels
VedGrabControl | 1 | 0=False/1=True | Enable/Disable frame grab for use with grab count averaging
GrabCount | 3 | Frames | Set the number of frames that the Grab Control will use when averaging frames. Averaging occurs on image freeze.
PointCloudSizeMin | 0 | Points | Sets the minimum number of points to be collected with the Measure Logic single fire probe
PointCloudSizeMax | 10000 | Points | Sets the maximum number of points to be collected with the Measure Logic single fire probe
IgnoreFovCalErrors | 0 | 0=False/1=True | *
AfAssist | 0 | 0=False/1=True | Enable or Disable the Focus Assist bar graphic(displayed with crosshair probes)
MinFocusRed | 255 | 0 to 255 | Focus bar histogram values above the value will be drawn red
MinFocusYellow | 0 | 0 to 255 | Focus bar histogram values above the value will be drawn yellow
CamSN | SNBuff | Serial Number | Specify Serial Number of IDS camera to be loaded
FrameHeight | 0 | Pixels | Sets the Height of the video frame
FrameWidth | 0 | Pixels | Sets the Width of the video frame
FOVNominalColor | 0, 255, 0 | RGB | Sets the color of the nominal lines in a profile feature (default is yellow)
FOVErrorColor | 255,0,0 | RGB | Sets the color of the error lines in a profile feature (default is red)
DebugFireDraw | 0 | 0=off, 1=on | Displays the location of data cloud points captured by VED by drawing perpendicular lines when enabled
DebugFireDelay | 40 | Ms | Delay between draw of each firing
RGBFilter | 0,0,0 | RGB Toggle 0=off, 1=on | Color filter function to improve edge detection for systems with monochrome bottom illumination in the RGB scale. Particulary when the color occupies one band predominantly.
Washout | 0 | 0=off, 1=on | Pixel saturation display mechanism. Off by default.
MotionUniformHist | 1 | 0=off, 1=on | Alternate motion detection algorithm for Auto Play mechanism. Disables motion detection attempt that uses Histogram of image, and uses edge based part detection instead. Histogram mode is the default.
AllowDxfFreeze | 0 | 0=off, 1=on | Surpress image freeze when translating or rotating DXF overlay files.
ShowIllumination | 6 | Number of Sample Zones (illumination tool) | Sets grid size in XY of illumination Sample Zones in addition to center zone & FOV boundary zones that are automatically included. Grid of samples only appears when using Cross Hair or Active Cross Hair
IlluminationRadius | 25 | Pixels (illumination tool) | Sets the Radius Size in pixels of each sample zone.
DefaultEyeMeasureWidth | 40 | Pixels | Set the default width of the EyeMeasure Probe
Flick | 0 | Enable=1, Disable=0 | Enable and disable of the touchscreen flick CNC functionality
CamSN_1 | 256 | Serial Number | IDS camera Serial Number. Index number is used in the M3 mag setup screen
PersistentOverlay | 0 | 0=off, 1=on | Overlay will remain in the Live Video view even after selecting different probes from the probe menu. Allows use of the simple crosshair for manual measurements while Overlay is visible
HideBrightFrameCal | 0 | 0=Show, 1=Hide | Allow the button in the Video Setup screen to be hidden
CamShareConfig | 0 | 0=off, 1=on | support for multiple IDS camera configuration files ueyecam_X.ini where “X” is the file index number to be assigned in the M3 softwares mag Setup Screen
ArchiveScale | 100 |  | Parameter for modifying the size of Full Frame Archive images exported from M3
CorrectDxfPixels | 1 | 0=False/1=True | Added pixel correction to the DXF Overlay based on FOV calibration data
[Model]
MinP | 85 | % | Sets the minumum required percentage of keypoint matches required for sucessfull pattern recognition
ClusterSizeFrame | 10 | Keypoint cloud cluster size | Parameter to adjust keypoint cloud cluster size used when searching for models
FilterMultiPatternROI | 0 | 0=False/1=True | Parameter to improve detection of multiple parts in the FOV when mutiple model matches exist for the same part.
[FOV]
DisplayDro | 1 | 0=False/1=True | Display or hide the Current Position DRO(For FOV systems)
[Profile]
IsOneSided | 0 | 0=False/1=True | Changes the way +P and -P coefficients are reported for profiles when all whiskers are on one side of the nominal. When parameter is set to 1 the +P and -P will report values of whiskers of least and greatest deviation from nominal, even when they are on the same side of nominal. When set to "0" the P from the side with no whiskers will report "0".
FilterAfterRun | 1 | 0=False/1=True | Enables a button to the bottom toolbar for removing fit points from a profile measurement after program playback. Fit results are recalculated immediately after whisker removal.
BlobRes | 1 | Integer | Sets the resolution of the Profile line segments created from the Blob points
DeleteImportedProfiles | 0 | 0=False/1=True | Governs profile deletion when constructing an M3 profile from an imported profile. By default imported profiles that are not consumed in a profile construction are automatically deleted. Enabling the ini below prevents the profile deletion.
[Registration]
AllowRotation | 1 | 0=False/1=True | Enable or disable rotation of the coordinate system in playback (beyond 45 degrees)
[Shortcuts]
On | 1 | 0=False/1=True | Create=K, Point=P, Line=L, Circle=C, Distance=D, Angle=A, Slot=S, Rectangle=R, Datum=Z, Skew=W If Z axis: Plane=E If 3d: Cone=N, Cylinder=Y, Sphere=H If digital comparator: Profile=F If show blob: Blob=B If v3: Stats=T If [Hardware]ZeroOnInputs=1: ZeroDro1=1, ZeroDro2=2 ,ZeroDro3=3, ZeroDro4=4 Y= button answers Yes and N= No when at a Yes/No Prompt. R= Run(Play) Done=<Enter> Quit=<Esc>
F1 | (none) | See above
F2 | (none) | See above | Send rs232 using function keys
F3 | R | See above | "
F4 | (none) | See above | "
F5 | X | See above | "
F6 | Y | See above | "
F7 | Z | See above | "
F8 | XY | See above | "
F9 | L | See above | "
F10 | D | See above | "
F11 | A | See above | "
F12 | XYZLDA | See above | "
[EmulatedStage]
ZXScale | 0.01 | Scalar | For every 100 x counts, the z is increaed by 1 counts (default values)
ZYScale | 0.05 | Scalar | For every 100 y counts, the z is increaed by 5 counts (default values)
[Optimize]
MaxStepsInARowBeforeUpdate | 10 | Program steps | Sets the number pf program steps that get updated before updating the results display
APFeaturesThreshold | 1000 | Number of Features | If more features then this are auto programmed, certain restricts (like max point count) are used to allow for reasonable program execution
UseNominalPtsAsActual | 0 | 0=False/1=True | Will playback the program using the nominal points for the measured points (debug tool)
[Debug]
AllowPaintLog | 2 | * | Enables main loop timer statistics (in the log)
[Window]
Left | * | Pixels | Sets the Application Window Pad from Left
Top | * | Pixels | Sets the Application Window Pad from Top
Width | * | Pixels | Sets the Application Window Width
Height | * | Pixels | Sets the Application Window Height
[MxCNCoptions]
UnitType | 0 | 0=As Stored, 1=Inch, 2=MM | Parameter to set the startup units for the mxcncoptions program.
ZoomType | 0 | Zoom Type | 0=Servo Zoom, 1=Opto Turrett, 3=GPIO Olympus Turret, 4=NikonTurret
DontKillLamps | 0 | 0=False/1=True
ReturnToHomePos | 1 | 0=False/1=True
[Zoom]
TurretRetractMM | 0 | MM | Amount that the Olympus Turret retract between mag changes
SkipMagPrompt | 0 | 0=False/1=True | Enables/Disables thedisplay of mag change prompt in program playback.
[Runs]
DeleteRunFiles | 0 | 0=False/1=True | When enabled, Run data files will be deleted from the PC harddrive, as they are removed from the M3 run view.
[Mlxaxis]
Spi_multiple | 2 | * | Used to slow down comunications to a 2 axis board. Faster pcs require a bigger number
[Demo]
EmulatedVelocity | 2 | Speed | Emulated Stage speed control parameter. The velocity of Goto moves within the emulated measuring stage can now be adjusted.
[Remote]
RemoteControlPortNumber | 3000 | Port Number | Added the ability to change the TCP port number for the "Remote Control" module, in the event of conflicts.